## Surname				First name			student’s code

# Section 1. TRUE/FALSE QUESTIONS

- The sensor span is the amplitude of the measuring range: 
	true o	false o

- The sensitivity threshold of a sensor is the resolution at the value 0 of the measuring variable: 
	true o	false o

- Closed liquid column manometers (“U” tube manometers) containing mercury are used for measuring vacuum 
	true o	false o

- The Coriolis Flowmeter is a transducer of volumetric flow rate
	true o	false o

- The hydrostatic sensor level does not require a float 
	true o	false o

- In gate valve the closure member always moves parallel to the seat
	true o	false o

# Section 2. MULTIPLE CHOICE QUESTIONS

- This sensor property depends on the full scale.

NOTE: check only the wrong answer!

- 	accuracy

- 	repeatability

- 	rangeability

- 	measuring range

- Which of the following types of level sensors is not used for liquids?

- 	conductive sensor

- 	vibrating sensor

- 	radar sensor

- 	probe sensor

- A magnetic flow meter determines the flow of a process material by measuring the _______ of the fluid

- 	velocity

- 	density

- 	temperature

- 	volume

# Section 3: SENSORS AND INSTRUMENTATION FOR PROCESS MEASUREMENTS

## 3.1.	Temperature measurement with thermocouple

A J-type thermocouple is used to measure the temperature of superheated steam.

The cold junction of the measuring circuit is at a temperature: Tcj = 305.16 K.

The voltmeter connected to the cold junction provides a measure of the electromotive force: V = 29.08 mV.

- determine the temperature Thj of the steam in °C with an approximation  1 °C

## 3.2.	Strain gauge (pressure transducer)

- Describe its working principle

## 3.3.	“Sonar” level sensor

- Describe its working principle

## 3.4 Rotameter

- Draw a sketch of its general working principle

- Derive its flow rate equation

- Extend the flow rate equation assuming P and T different from calibration condition.

- Discuss briefly advantages e disadvantages

NOTE: A well organized and short text will be assessed more than a long and confused one!

# Section 4: VALVES

## 4.1 Valves

Draw a simple sketch, explain purpose and features of each one of the following valves.

- Ball cock

- Butterfly control valve

- Angle valve

- Three-way valve

NOTE: A well organized and short text will be assessed more than a long and confused one!

## 4.2. Sizing Problem

A control globe valve has to be chosen and sized at the following conditions:

fluid: acetic acid

density: f = 8.75 lb/gal (US)

nominal flow rate: m = 15 lb/s

nominal pipe size: NPS = 120 mm

upstream pressure: P1 = 3 atm

downstream pressure: P2 = 1.7 atm

vapor pressure: Pv = 1.276 kPa

liquid critical pressure ratio factor: FF = 0.956

- Calculate the flow coefficient Cv

A steel-made, flanged-mounted globe valve is available. It may adopt either a linear or equalpercentage or parabolic inherent characteristic. The rangeability is r = 50. The valve manufacturer data are:

| NPS mm | Kvn m3 (H2O) /h bar½ |
| --- | --- |
| 8 | 3.0 |
| 15 | 9.2 |
| 20 | 12.1 |
| 40 | 17.0 |
| 60 | 29.3 |
| 80 | 34.6 |
| 100 | 52.8 |
| 120 | 70.3 |
| 150 | 88.4 |

The conversion for Kv to the flow coefficient Cv is as follows:

Cv = 1.16 Kv [=] gpm (USA) psi-1/2

- What is the installed characteristic?

If the Combraco 57 valve is already installed in a circuit plant:

- Choose an inherent characteristic that would be congruent to the following value of the valve authority V = 0.2

- Size the valve by selecting the proper NPS as a consequence of the selected  inherent characteristic

- Check cavitation according to the IEC norm

- Calculate the value of the vapor pressure Pv for which cavitation occurs in the valve

- Choose an inherent characteristic that would be congruent to the following value of the valve authority V = 0.5

- Size the valve by selecting the proper NPS as a consequence of the selected  inherent characteristic

- Do the valves sized under the items 3) and 7) have the same NPS? Why?